POWDER ED-XRF ANALYSIS PROTOCOL

By Martin P. Pentrak, Ph.D.,

XRD & XRF Lab Manager

Illinois State Geological Survey, Prairie Research Institute, University of Illinois at Urbana-Champaign

Sample preparation and analyses were carried out at the X-Ray Fluorescence Laboratory at the Illinois State Geological Survey (ISGS) Prairie Research Institute, University of Illinois at Urbana-Champaign.

METHODOLOGY

Sample preparation

The whole samples are dried at 105°C overnight, gently ground, sieved, and less than 250 µm fraction is collected.

The XRF analysis of the powder samples

Powder samples are analyzed by Rigaku NEX CG Energy-Dispersive X-Ray Fluorescence (ED-XRF) spectrometer using a non-destructive methodology. Samples are placed into sample holders fitted with a Chemplex Industries, Inc. Prolene® films, then placed into the spectrometer sample chamber equipped with 15-position sample changer above the X-ray source window and measured under an He atmosphere. The instrument specifications include a 50-watt Pd-anode end window X-ray tube, four secondary targets (RX9, Cu, Mo, and Al), a 20 mm diaphragm, and a high-performance silicon drift detector (SDD). Secondary targets with measuring time 100 s (RX9, Cu, and Mo) and 200 s (Al) provide high peak to background (P/B) ratios compared to direct excitation instruments.

The XRF data analysis

Semi-quantitative analyses are obtained using Rigaku’s RPF-SQX (Rigaku Profile Fitting-Spectra Quant X) software combined with the Powder FP (fundamental parameters) method. The software determines the lower limit of detection (LLD or LOD) defined by the lowest concentration distinguished from the background. Precision and accuracy are determined by analyzing sample triplicates as well as certified references materials from the National Institute of Standards & Technology (NIST).