POWDER XRD ANALYSIS PROTOCOL

By Martin P. Pentrak, Ph.D.,

XRD & XRF Lab Manager

Illinois State Geological Survey, Prairie Research Institute, University of Illinois

The X-ray diffraction analysis of randomly oriented powder samples (mineral phases identification and quantification) is performed in the XRD laboratory (Martin P. Pentrak – Lab Manager) at the Illinois State Geological Survey (ISGS) within the Geochemistry Section, Prairie Research Institute, University of Illinois at Urbana-Champaign, IL, USA.

METHODOLOGY

Sample preparation

The whole samples are dried at 105°C overnight, gently ground, sieved, and less than 250 µm fraction is collected.

The XRD analysis of the bulk powder (XBP)

5 grams of prepared powdered (< 250 μm) samples are further gently ground by mortar and pestle and then the fine powders are packed into an end-loading sample holder as a random powder bulk-pack. The powder samples are analyzed with a Scintag® XDS2000 XRD instrument with a step-scanned data collected from 3° to 70° 2θ (random powder bulk-pack) with a fixed rate of 1° per minute with a step size of 0.02°2θ for each sample.

The XRD data analysis

The powder XRD (XBP) analysis provides a qualitative (mineral phase identification) analysis based on the location and intensity of peaks of the 2θ scale compared with the powder diffraction files (PDF2) database. The quantitative analysis (samples spiked with an internal standard: 20% wt. Al2O3) is performed by the Rietveld refinement procedure using data reduction software from Materials Data Inc. (MDI) known as Jade+®.